| STAR5000C Patient Monitor |  |  |
| --- | --- | --- |
| Safety Standards |  |  |
| ISO 13485:2003 approved, CE marking according to MDD93/42/EEC |  |  |
| Physical Characteristics |  |  |
| Product Size: | 420㎜×250mm×340mm |  |
| Net Weight: | 5.0kg |  |
| Display: | 12.1’’ Color TFT LCD touch screen |  |
| Resolution: | 800X600 |  |
| Method: | folded |  |
| Operation Environment |  |  |
| Temperature: |  | 0-40℃ |
| Humidity: |  | 15-85% |
| Power Supply |  | AC100-250V, 50/60Hz |
| Battery Type: |  | Rechargeable lithium-ion battery |
| Battery Capacity: |  |  |
| Battery Recharging Time: |  | Maximum 6.5 hours for charging; |
| Battery working period: |  | 2 hours for continuous working |
| Interface: |  | USB(data input or output, upgrade) AC power supply socket Parameter cable interface RS232 port RJ45 port |
| Data: |  | Input/output(patient’s case) |
| Network: |  | Central Monitor System |
| Performance Specifications |  |  |
| Printing |  |  |
| Recording parper: |  | 112mm*90mm*150 Z-fold, thermo sensitive |
| Fast print speed: |  | 25mm/s(stored traces) |
| Paper Advance Speed: |  | 1/2/3cm/min selectable |
| Record Message: |  | Data,t ime, TOCO type,paper speed,FHR type,bed no.,etc. |
| Printer: |  | Built-in thermal recorder; external laser printer |
| Ultrasound |  |  |
| Technique: |  | Ultrasound Pluse Doppler with autocorrelation |
| Pulse repetition rate: |  | 2 KHz |
| Ultrasound: |  | 1.0MHz ±10% |
| FHR Range: |  | 30bpm- 240bpm |
| Resolution: |  | 1bpm |
| Accuracy: |  | ± 2 bpm |
| TOCO |  |  |
| TOCO range: |  | 0-100 Relative(%) |
| Resolution: |  | 1% |
| Zero Mode: |  | Manual |
| Non-linear error: |  | ≤±10% |
| Fetal Movement(FM) |  |  |
| Mode: |  | Auto; manual |
| Technique: |  | Pulsed Doppler Ultrasound |
| Range: |  | 0-100(%) |
| Resolution: |  | 1% |
| Fetal Stimulator |  |  |
| Method: |  | Vibration |
| Mode: |  | Manual(3 times), continuing |
| Voltage: |  | DC 3V, (AA battery) |
| Respiration |  |  |
| Method: |  | impedance method |
| RR measurement range: |  | Adult: 0-120bpm Neonate:0-150bpm |
| Resolution: |  | 1 rpm |
| RESP Apnea Alarm: |  | 10s-40s |
| Alarm: |  | Audible and visual alarm; alarm events recallable |
| Sweep Speed: |  | 6.25、12.5、25mm/s |
| Gain Selection: |  | X1, X2, X4 |
| ECG |  |  |
| Lead Type |  | 5 Lead and 3 lead selectable |
| 5-lead cable: |  | RA; LA; RL; LL; V or R; L; N; F; C |
| 3-lead cable: |  | RA; LA; LL or R; L; F |
| Lead selection: |  | 5-lead: I; II; III; aVR; aVL 3-lead: I; II; III |
| Gain Selection |  | X0.25, X0.5, X1, X2 |
| Sweep Speed |  | 12.5, 25, 50mm/s |
| Heart Rate Range |  | Adult: 15-300bpm Pediatric/Neonate:15-350bpm |
| Resolution: |  | 1 bpm |
| Protection: |  | Withstand 4000VAC/50Hz voltage in isolation; Against electrosurgical interference and defibrillation; |
| Accuracy: |  | ±1% or ±2bpm (whichever is greater) |
| Filter |  | ECG enlarge width: MON: 0.5-55Hz DIA: 0.05-150Hz OPE:1-20Hz |
| ST SEGMENT detection: |  | -2.0mV - +2.0Mv (Automatic) |
| Arrhythmia analysis : |  | Yes |
| pacemaker detection: |  | detectable |
| NIBP |  |  |
| Method | Automatic oscillometric |  |
| Work mode: | Manual / Automatic/Continuous |  |
| Auto measure time: | Adjustable(1-480min) |  |
| Measurement Unit: | mmHg / Kpa selectable |  |
| Measurement types: | Systolic, Diastolic, Mean |  |
| Range of systolic pressure: | Adult Mode:40-270mmHg Pediatric Mode:40-200mmHg Neonate Mode 40-135mmHg |  |
| Range of diastolic pressure: | Adult Mode:10-215mmHg Pediatric Mode:10-150mmHg Neonate Mode 10-100mmHg |  |
| Range of mean pressure: | Adult Mode:20-235mmHg Pediatric Mode:20-165mmHg Neonate Mode 20-110mmHg |  |
| Over-pressure protection: | both Hardware and software over pressure protection |  |
| Accuracy: | Less than 5mmHg |  |
| Resolution | 1mmHg |  |
| Alarm | Systolic, Diastolic, Mean |  |
| PR from NIBP: | 40-240bpm |  |
| Temperature (Dual Channel) |  |  |
| Measurement range: | 0-50C |  |
| TEMP sensor: | Standard configuration- skin TEMP sensor, |  |
| Resolution: 0.1C |  |  |
| Accuracy: | 0.1C (exclusive of error of sensor) |  |
| Channel: | T1, T2, TD (Temperature Difference) |  |
| Comen SPO2 |  |  |
| Measurement & alarm range | 1~100% |  |
| Resolution: | 1% |  |
| Accuracy: | ±2% (70~100%,Ped, non-motion+ ±3% (70-100%, Neo, non-motion); 0-69% unspecified |  |
| Alarm range | 0~100% |  |
| Pulse rate: | Range: 20~250bpm |  |
|  | Resolution:1bpm |  |
|  | Accuracy: ±3bpm(non-motion) |  |
|  | ±5bpm(motion) |  |
|  | Alarm range:20~250bpm |  |

Standard Configuration :

Single FHR,TOCO,FM,NIBP,SPO2,ECG,RESP,PR

Optional:

Twins , fetal stimulator