General Information:

| Product Description | Wrist Blood Pressure Monitor |
| --- | --- |
| Sales Name | RS1 |
| Item number | HEM 6120 E |
| Model Category | Wrist Type Blood Pressure Monitor |
| Legal Manufacturer |  |
| Production Facilities | OMRON (DALIAN) Co., LTD. Dalian, China |
| Photo of Product |  |
| Packaging Content | Main unit, pre-formed cuff, battery pack (2 x AAA), Instruction manual, Guarantee card |
| EAN Code | 4015672107274 |

Packaging Information:

| Carton | Qty | Approx Weight | Approx dimensions W x D x H mm |
| --- | --- | --- | --- |
| Main unit without batteries | 1 | 101g | 78 (w) mm x 60 (h) mm x 21 (d) mm (without the wrist cuff) |
| Sales Package | 1 | Tbd – after first mass production | H110mm x W 100mm x D110mm |
| Master Carton | 20 | Tbd – after first mass production | tbd |

Particularities

RS1 (HEM 6120 E) is an automatic wrist type blood pressure monitor that has the unique selling points as noted:

•	Intellisense Technology

•	Cuff wrapping guide

•	Optimal speed and pressure for gentle inflation/deflation

•	Pre-formed easy to fit cuff

•	Last memory

TECHNICAL PRODUCT INFORMATION

| Product Description | Digital Automatic Wrist Blood Pressure Monitor |
| --- | --- |
| Model | OMRON RS1 HEM-6120-E |
| Display | LCD Digital Display |
| Measurement Method | Oscillometric method |
| Languages | EN, DE, FR, IT, ES, NL, RU, AR, |
| Measurement Range | Pressure: 0mmHg to 299mmHg Pulse: 40 to 180/min. |
| Memory | Last Memory |
| Applied Part | = Type B |
| Protection Against Electric Shock | Internally powered ME equipment |
| Accuracy | Pressure: ±3 mmHg Pulse: ±5% of display reading |
| Inflation | Automatic inflation by pump |
| Deflation | Automatic rapid deflation |
| Pressure Detection | Capacitive pressure sensor |
| Power Source | 2 “AAA” alkaline (LR03) batteries 1.5V |
| Battery life | Approximately 300 measurements with new alkaline batteries at a room temperature of 23°C |
| Operating temperature / Humidity | +10C to +40C Maximum: 30% to 85% RH |
| Storage Temperature / humidity | -20C to +60C Maximum: 10% to 95% RH |
| Console Weight | Approximately 101g without batteries |
| Outer dimensions | Approximately 78 (w) mm x 60 (h) mm x 21 (d) mm (without the wrist cuff) |
| Measurable circumference of wrist | Approximately 13.5 to 21.5 cm |
| Cuff material | Nylon and polyester |
| Package Content | Main unit, battery set, instruction manual, guarantee card |

This device fulfils the provisions of EC directive 93/42/EEC (Medical Device Directive).

• This blood pressure monitor is designed according to the European Standard EN1060, Noninvasive

sphygmomanometers Part 1: General Requirements and Part 3: Supplementary

requirements for electromechanical blood pressure measuring systems.

• This OMRON product is produced under the strict quality system of OMRON HEALTHCARE

Co. Ltd., Japan. The Core component for OMRON blood pressure monitors, which is the

Pressure Sensor, is produced in Japan.

Intended Use

Medical Purpose

This product is intended for use to measure and/or monitor patients to record their blood pressure levels.

Intended User

• Legally certified medical experts, such as doctor, nurse and therapist, or healthcare personnel or patient under the guidance of qualified medical experts.

• The user should also be capable of understanding general operation of this device and the content of instruction manual.

Intended Patients

This product should not be used by patients, who are unconscious or are not breathing spontaneously, unattended children or persons who cannot express their intent.

Environment

This product is intended for use in a medical facility, clinic and doctor’s office, and in a room of general household.

Packaging

Expanded View – Front Panel

Expanded View – Back Panel

Expanded View - Side Panel 1

Expanded View – Side Panel 2

Comparison Table – there was no previous E version

Instruction Manual Design